Imaging Centre of Excellence, QEUH

7T MRI

LOCAL RULES &

OPERATING PROCEDURES

Signature of MR Responsible Person:		Date: 01/03/21

Signature of MR Safety Expert:	   		Date: 03/03/21

This document adheres the MHRA Safety Guidelines for Magnetic Resonance Imaging Equipment in Clinical Use. February 2021. The latest version can always be found at:

https://assets.publishing.service.gov.uk/government/uploads/system/uploads/attachment_data/file/958486/MRI_guidance_2021-4-03c.pdf

Index

- Introduction

- The hazards in MRI

- Responsibilities

- Control of Access

- Subject management

- Categories of Staff and Essential Training

- Equipment

- Emergency Procedures

- Approved implants, medical  devices and equipment for the 7T environment

Appendices

Appendix one		Plan of facility with magnet and 5 Gauss line

Appendix two		Safety labelling

Appendix three	Checklists

- Introduction

  - Purpose

The aim of these Rules and Operating Procedures is to provide advice and guidelines for the safe use of the Magnetic Resonance Equipment with a particular view to ensuring the safety and well being of subjects, staff and all personnel involved in any respect or likely to be so involved with the MR Imaging Facility.  They are also intended to ensure that there is at all times compliance with the responsibilities and requirements of the Health and Safety at Work Act 1974.

## Background

The equipment installed is a Siemens Terra 7T system. During MRI, individuals being scanned are exposed to static magnetic, time-varying magnetic and radiofrequency (RF) fields.

In 2004 the International Commission on Non-Ionizing Radiation Protection (ICNIRP) produced guidelines on the protection of MRI subjects [ICNIRP 2004].  In the light of the increasing number of ultra-high field (7T and above) MRI systems being installed worldwide, ICNIRP reviewed the effects of ultra-high fields [ICNIRP 2009] and issued an amendment to their previous guidelines [ICNRIP 2009a].  Both the 2009 reports make reference to the ICNIRP guidelines on limiting exposure to time-varying electrical and magnetic fields (revised 2010) [ICNIRP 2010].

The UK Health Protection Agency (HPA) and the Medicines and Healthcare products Regulatory Agency (MRHA) both relied heavily on ICNIRP guidelines to produce their documents on the protection of subjects and volunteers undergoing MRI procedures [HPA 2008], static magnetic fields [HPA 2008a] and guidelines for MRI equipment in clinical use [MHRA 2011].

In the USA, the Food and Drug Administration (FDA) have also addressed this issue [FDA 2003].

## References

The main reference documents for these Rules and Operating Procedures are:

- Safety Guidelines for MR Equipment in Clinical Use:  MHRA, March 2015

- Guidelines for Magnetic Resonance Equipment in Clinical Use: issued by the UK. Dept of Health, February 1993.

- The Control of Electromagnetic Fields at Work Regulations, Health and Safety Executive, UK, 2016

  - Definitions used in Local Rules

The defined terms in this document as recommended by the MHRA March 2105 are as follows:

1.4.1    MR Controlled Access Area

A locally defined volume of such a size to contain the MR Environment, see appendix one for a plan of the facility. Access should be restricted and suitable signs should be displayed at all entrances.

1.4.2     MR Environment

The volume within the 0.5 mT line of an MR system, which includes the entire three-dimensional volume of space surrounding the MR scanner. For cases where the 0.5 mT line is contained within the Faraday shielded volume, the entire room shall be considered the MR environment. This volume is the region in which an item might pose a hazard from exposure to the electromagnetic fields produced by the MR equipment and accessories.

Important Note: For the purpose of this document all safety precautions will apply to the MR Environment, which includes both the 5 Gauss and 30 Gauss contours. This includes the magnet room of the scanner plus the equipment / computer room. The Controlled Access Area referred to will be any area within the Department, beyond the secure entrance.

    - MR Safe

An item which poses no known hazards in all MR environments

    - MR Conditional

An item which has been demonstrated to pose no known hazards in the specified MR environment within defined conditions which include field strength, spatial gradient, dB/dt (time rate of change of the magnetic field), radio frequency (RF) fields and specific absorption rate (SAR). Additional conditions, including specific configurations of the item may be required.

    - MR Unsafe

An item which is poses unacceptable hazards within the MR environment.

1.4.6     MR Authorised Person

An MR authorised person is a suitably trained member of staff authorised to have access to the MR controlled access area.

- Authorised Person (Non-MR Environment): An MR authorised person authorised to have free access to the MR controlled access area but not the MR environment. They may supervise other persons only in this area.

- Authorised Person (MR Environment): An MR authorised person authorised to have free access to the MR environment but not to supervise others.

- Authorised Person (Supervisor): An MR authorised person who is authorised to have free access and to supervise others in the MR environment.

A full list of current MR Authorised persons can be found on CRIF-MRP-ICE-001 https://gla-my.sharepoint.com/:x:/g/personal/tracey_steedman_glasgow_ac_uk/ESAFLQjcJUpAjGZiRuH95tcB_VslKUB0pM7hzZdaGu05mw?e=8BSE3q

1.4.7    MR Operator

An MR operator is an MR authorised person who is also entitled to operate the MRI equipment.

The Terra system has two operating modes – single transmit (sTX) and parallel transmit (pTx) therefore operators may only work with the mode they have been authorised for.

1.4.8     MR Responsible Person

It is recommended that the chief executive or the general manager delegate the day-to-day responsibility for MR safety to a specified MR responsible person who might most effectively be the clinical director, head of the department, clinical scientist, medical physicist or MR Superintendent Radiographer of the institution where the equipment is located.

1.4.9     MR Safety Expert

The MR safety expert will have an advanced knowledge of MRI techniques and an appropriate understanding of the clinical applications of MRI. Ideally they will be a physicist with expertise in MRI.

- Hazards of MRI

  - Static Magnetic Fields (B0) Related Issues

The Siemens Terra System is built around a superconducting magnet with field strength of 7 Tesla.  This magnet is always on, even if no scanning is taking place.

MRI researchers are trained in and familiar with the safety issues related to working with strong magnetic fields and, although there are no specific additional safety issues involved in working at 7 Tesla, above say 3 Tesla, the magnitude of the forces involved and effects experienced are greater.

Safety issues to consider with a strong static field Bo are: biological effects, projectile hazards, compatibility of implantable medical devices and compatibility of peripheral equipment.

    - Fringe Fields

All staff who works in MRI must be well informed about fringe fields and should know the location of the 0.5 mT (5 Gauss) and 3 mT (30 Gauss) contours. Appendix 1 illustrates the plot of the Terra system. Note that the 0.5mT does not extend into the corridor nor control room. It does however extend into the Technical Room and so this room should always be clearly marked.

    - Biological Effects

Safety documentation   reports that subjects exposed to magnetic fields of up to 8T sometimes experience effects of:

- Nausea

- Dizziness

- Phosphenes (visual phenomena, usually flashes of light)

- Metallic taste in the mouth

These effects were transitory and no long-term exposure effects were detected. These effects are strongly related to movement through the changing field (from 0 to 7 T), rather than field strength per se, and can be reduced by moving slowly through the field contours as the subject is inserted into and removed from the scanner.

Since the field changes associated with the 7T system are steeper than our current 3T scanners we would expect a greater proportion of scan volunteers to feel these effects.

However experience from other 7T sites, including Nottingham University, suggest that most subjects experience only a sensation of ‘going round a corner’ as they are moved slowly into the magnet (despite the magnet bore being straight).

    - Attractive Force

Not all metals are ferromagnetic (attracted by a magnetic field) but for those that are, a strong attraction force exists.  Ferromagnetic objects may pose projectile risks, while ferromagnetic implanted devices may malfunction in a strong magnetic field.

The forces involved depend on the size and the gradient of the static magnetic field and the size and shape of the object, but it is appropriate to assume that the forces on these objects will be stronger at 7 Tesla, compared to 3 Tesla.  The force experienced in the Terra system is at its maximum just inside the bore as illustrated in figure 1:

Figure 1 – Cross-sectional field plot of the Terra system.

The types of ferromagnetic objects that a subject may have on their person will either be removable objects, such as keys, coins, watches, belts or implanted objects such as surgical pins, pacemakers or stimulators.  Many such implanted objects have been tested for safety at 3 Tesla, but most will not have been specifically tested at 7 Tesla.  However, since the mechanism of the force is the same at all field strengths, an experienced MR physicist or radiographer would be able to extrapolate 3T risk or lack of risk to 7 Tesla.

2.1.4  Control of Equipment Taken Into the MR Environment

All peripheral equipment suitable for taking into the MR Environment must be clearly and suitably labelled. Refer to section 1.4.3-5 for definitions. Appendix two illustrates appropriate labelling.

Any equipment deemed to be non-ferrous will be marked using the MR safe label. Only equipment so marked may be taken into the MR Environment by Authorised Personnel. It is the shared responsibility of the MR Responsible Person and MR Safety Expert to ensure that items are regularly marked and checked. Any devices or equipment not labelled appropriately must be considered MR Unsafe.

Control of equipment entering the MR ENVIRONMENT on a day-to-day basis is the responsibility of the MR OPERATOR responsible for the examination at the time.

    - Torque

Ferromagnetic objects will also experience a torque that will try to align that object along magnetic field lines. Torque is largely shape-dependant and is proportional to the field strength, B, and to the angle the object is away from alignment with the field.

    - Interaction with Implantable Medical Devices

When a conductor moves through the flux of a magnetic field, a potential difference is induced that is proportional to the rate of change of the flux. Lenz’s Law states that the induced potential difference is in a direction to oppose the change inducing it. The result is to induce a magnetic field in the moving conductor, which will resist that movement.

Effects are minimal up to 1.5 T but can be significant at 7T. This effect could reduce the working of mechanical valves, such as mitral and aortic valve replacements.

Only items listed in the Approved Implants folder (https://gla-my.sharepoint.com/:f:/g/personal/tracey_steedman_glasgow_ac_uk/EqmwSxAjIdBHsqoH7DBch-EBXGQvYQgZyhqygsvZqEFOUw?e=ArZBV6) may be scanned.

    - Interaction with Ancillary Equipment

The static magnetic field can affect monitoring equipment that has ferromagnetic components.

This could affect the function of the equipment, due to both magnetic field and currents induced by movement through a static magnetic field and / or cause a projectile hazard.

It is recommended that only monitoring equipment intended for use in a 7T MR Environment be used.

Accessories to monitoring equipment should also be checked for compatibility e.g. ECG electrodes.

  - Time-Varying Magnetic Field Gradients (dB/dt)

All MRI scanners contain electromagnetic coils that generate time and spatially varying magnetic fields. The Siemens Terra 7 T scanner is equipped with a gradient coil that has a peak magnetic field gradient of 80 mT/m and a peak slew rate (rate of change of gradient) 200T/m/s on each axis. This is the same as the Siemens Prisma 3T and so there is no further risk.

Safety concerns with the time-varying magnetic field gradients are biological effects, peripheral nerve stimulation, muscle stimulation and acoustic noise.

In MR, three orthogonal magnetic field gradients are switched on and off to select the region of diagnostic interest and to spatially encode the MR signals. As a general guide the faster the sequence, the greater the rate of change of the gradient fields used and the resultant current density induced in the tissues.

    - Biological Effects

Time-varying magnetic field gradients can lead to induced electric fields and circulating currents in conductive tissues. Induced electric currents can interfere with the normal function of nerve cells and muscle fibres e.g. causing peripheral nerve stimulation. A more serious response is ventricular fibrillation, which is prevented when the Terra is operated within IEC limits i.e. in single channel mode.

    - Peripheral Nerve and Muscle Stimulation

At low frequencies, induced currents are able to produce the effect of stimulation of nerve and muscle cells. The extent will depend on the pulse shape and its repetition rate.

It can be sufficient to cause discomfort and in extreme cases might result in limb movement or ventricular fibrillation. There have been reported incidents of subjects and volunteers experiencing PNS whilst undergoing MR examinations.

  - Radiofrequency Magnetic Fields (B1)

The main safety issues for radiofrequency (RF) fields used in MR are thermal heating leading to heat stress induced current burns and contact burns. Absorption of energy from RF fields results in the increased oscillation of molecules and the generation of heat.

Heat is dissipated through the skin. Organs with reduced blood flow take longer to dissipate heat. 1oC is acceptable for normal healthy individuals. The lower the ambient temperature and the lower the humidity, the greater the transfer.

    - Modes of operation and SAR

The Terra system has two operating modes, singe transmit (sTx) and parallel transmit (pTx).

Compliance with appropriate SAR restrictions will be provided automatically by the scanner in sTx mode.

However, when scanning in pTx mode this is not the case. It is therefore mandatory to record SAR values of each sequence during all protocols, every time, using CRIF-ICE-FORM-004 (https://gla-my.sharepoint.com/:w:/g/personal/tracey_steedman_glasgow_ac_uk/EWfTCgoabJNEuFCyBPj0UpQBsOwAylIvXWWRh3Xek6R80w?e=PP5EEx

2.3.2    Heat Stress

Heat stress is of particular concern for subjects who are hypertensive, pregnant or are on drugs such as diuretics or vasodilators. This subject group should be exposed to the minimum RF level possible.

    - Burns

###### Contact Burns:

- The RF field will induce currents in conductors and can raise their temperature significantly.

- Subject contact with any metallic object must be avoided by careful positioning of subject and coil.

Induced Current Burns:

- Caused by subject’s arms / legs being placed in such a way to create a conductive loop pathway.

To avoid both contact and induced current burns, use foam pads 1-2 cm thick to prevent subject from touching cables, the magnet bore and themselves.

    - Acoustic Noise

As gradient switching occurs, alternating low-frequency currents flow through the gradient coils, which are immersed in the high static magnetic field. Forces are then exerted on the gradient coils that move like a loudspeaker coil and generate sound waves.

Ear-plugs, ear defenders or other means of hearing protection must be administered to all individuals who are within the MR Environment during scanning.

Groups of particular concern are paediatric, neonates, the foetus, unconscious subjects and those with pre-existing aural conditions such as tinnitus, recruitment or hypersensitivity.

    - Pregnancy

Pregnant participants: Current literature indicates that no pregnant persons should be scanned at 7T.

Pregnant Staff: Pregnant staff will not enter the MR Environment during the first trimester of pregnancy. They will be risk assessed at regular intervals as required to assess health and wellbeing. Pregnant staff will not be permitted to enter the scan room whilst scanning is underway at any stage during the pregnancy due to the concerns of acoustic noise exposure and risks to foetus.

    - Cryogens

All scanners have a quench pipe that should vent helium and nitrogen directly to outside of building in case of boil-off or pressure release valve bursting.

Hazards in use of low temperature liquefied gases for MR systems are asphyxiation in oxygen-deficient atmospheres, cold burns, frostbite and hypothermia from intense cold and explosion following over-pressurisation from the large volume expansion of the liquid following evaporation.

## Responsibilities

For optimum safety to be achieved in any organisation there must be a joint understanding of the responsibilities of management and the responsibilities of individuals. Management and individuals must be fully aware, at all times, of the need for safety and the consequences that may arise if vigilance is relaxed.

Within the ICE building the NHS and therefore chief executive or general manager, has responsibility for all aspects, of safety with respect to the equipment, its location, its use, the subjects scanned, and all personnel who have access to the equipment location.

### 7T facility organisational structure

- UoG ICE Director				Dr Carol Clugston

- UoG Clinical Lead				Prof Keith Muir

- UoG Scientific Leads			Prof Lars Muckli/ Ms Jozien Goesne

- NHS GG&C R&D Director			Prof Julie Brittenden

- NHS GG&C MR Responsible Person	Mrs Tracey Hopkins

- NHS GG&C  MR Safety Expert		Dr John Foster

- A full list of current MR Authorised persons can be found in CRIF-MRP-ICE-001 https://gla-my.sharepoint.com/:x:/g/personal/tracey_steedman_glasgow_ac_uk/ESAFLQjcJUpAjGZiRuH95tcB_VslKUB0pM7hzZdaGu05mw?e=8BSE3q

### MR Responsible Person

The Lead Research Radiographer for NHS GG&C R&D, Mrs Tracey Hopkins, will act in the capacity of Responsible Person and have responsibility for:

- All aspects of safety within the Magnetic Resonance Imaging Facility

- Provide access to appropriately trained users

- For ensuring that Local Rules are reviewed and updated at regular intervals

- For ensuring that all staff involved at any time with the equipment or the Facility are appropriately trained.

- On a day-to-day basis, the Authorised MR Operator will act as the Agents of the Responsible Person.

  - MR Safety Expert

The NHS GG&C Safety Expert, Dr. John Foster, will act in the capacity of MR Expert providing scientific advice to the MR Responsible Person. They have responsibility for:

- Advising on the necessary engineering, scientific and administrative aspects of the safe clinical use of the MR devices including site planning, development of a safety framework, advising on the effectiveness of local safety procedures, procurement, adverse incident investigation and advising on specific participant examinations.

  - Designated Person

The Designated Person role ensures that there is always clarity of leadership in the MRI facility.

In the absence of the MRRP the designated duty CRIF Radiographer will act in this role. This will normally be the Deputy Lead Research Radiographer.

In the absence of the MRSE a delegated duty MRSE will act in as the designated person.

## Control of Access the MR Controlled Access Area

It is absolutely vital to control access of personnel and equipment to the MR Controlled Access Area and to control those individuals who are scanned.

  - Authorised Persons

An MR Authorised Person is a suitably trained member of staff authorised to have free access to the MR Controlled Area. This includes MR Operators and Siemens Engineers.

|  | AUTHORISED PERSONNEL MR ENVIRONMENT | MR CONTROLLED ACCESS AREA outside MR ENVIRONMENT. |
| --- | --- | --- |
| AUTHORISED PERSON (NON-MR ENVIRONMENT) | May not enter without supervision | May enter and supervise |
| AUTHORISED PERSON (MR ENVIRONMENT) | May enter | May enter and supervise |
| AUTHORISED PERSON (SUPERVISOR) | May enter and supervise | May enter and supervise |

Table 1: Categories of Authorised Persons

4.1.1   Authorisation of Personnel

Personnel become authorised for each scanner after completing the CRIF Training Schedule, which is signed off by the MR Responsible Person.

It is the responsibility of all Authorised Personnel to notify the MR Responsible Person / of any potential changes in their suitability to enter the MR environment.

    - Screening of MR Authorised Personnel

All MR Authorised Personnel must have satisfactorily passed the screening process, with repeat screening taking place annually. All MR Authorised Personnel must satisfy themselves that they conform at all times to the requirements of the screening process.

    - Responsibilities of MR Authorised Personnel

- They must comply at all times with the safety recommendations for the Controlled Access Area.

- Ensure any other persons who require access are accompanied by an MR Authorised person

- Take on full responsibility for the presence of the unauthorised person or persons for the duration of their presence in the MR Controlled Area.

    - Unauthorised Persons

ALL other persons which includes study participants, visitors, maintenance staff, domestics who enter the MR Controlled Access Area may only do so if:

- They have had verbal explanation of potential hazards within the MR Environment.

- They are supervised at all times by an MR Authorised Person.

In addition to the above if entering the MR Environment unauthorised persons must also:

- Have completed and signed a satisfactory safety screening form.

- Have removed all ferromagnetic materials

4.2    Categories of Staff and Essential Training

To avoid accidents, it is essential that all personnel associated with the MR equipment have been adequately trained and authorised by the MRRP.

The list below gives examples of typical categories of staff who will or may need to enter the MR Controlled Access Area in the course of their duties.

The choice of individuals, their number and their category needs careful consideration before authorisation is given.

- Category (A): MR Operator: Those wishing to operate, maintain or modify the MR equipment such as radiographers, scientific staff, technical staff and in certain cases suitably qualified and trained research associates and students.

All operators must be signed-off and authorised by the MRRP separately for sTx mode scanning and pTx mode scanning.

Operators wishing to utilise IDEA for sequence development on the Terra must be authorised to do so by Prof. David Porter.

- Category (B): Personnel who do not fall into category (A) but are present with a volunteer or subject during scanning such as clinicians and research assistants.

- Category (C): All staff who are required to enter the MR Environment when scanning is not taking place e.g. dedicated cleaning, estate maintenance staff. Supervised at all times.

- Category (D): All other staff who are required to enter the MR Controlled Access Area but will not enter the MR Environment.

Table 2: Summary of training requirements for authorised persons

### Security

Access to the MR Controlled Access Area will be through self-locking, card entry doors for which only Authorised Personnel are able to gain access. Relevant keys are kept in a locked key cupboard / drawer. All Authorised MR Staff must know the location of keys.  This information must not be divulged to non-authorised personnel.

Only the MR Responsible Person is able to enable card access to the MR Controlled Access Area once training has been completed.

The Magnet Door will remain open during the day, as long as there are Authorised Personnel in the Control Room. The Technical Room within the MR Environment is kept locked.

Any malfunction of a self-locking door must be reported immediately to the Authorised or Responsible Person. Under no circumstances must these self-locking doors be propped in the open position.

Outside the MR Environment, lockers are made available for valuables, ferromagnetic materials and magnetic recording media belonging to subjects and visitors.

  - Researcher Access Out of Hours

This is defined as working out with of the core working hours of 9-5pm, Monday to Friday. Out with these hours access is only gained with the NHSGGC ID card.

The ID card can only be issued by the MRRP once appropriate level of training has been completed and entitlement and sign off given (section 4).

A minimum of two MR authorised persons must be in the MR Controlled Access Area at all times to comply with MHRA Safety Guidelines. This must include a minimum of one MR Authorised Operator who will supervise the MR Authorised Non-operator

No access will be given to MR Environment under any circumstances unless by or under the Supervision of an MR Authorised Operator.

- Participant Management

All subjects must confirm their identity i.e. name (including spelling), date of birth and address during the screening process. Once confirmed the subject and MR Authorised Person must sign the MR Safety Screening Checklist.

  - Screening Prior to Examination

Screening is important to assess the suitability of a subject for MRI. This includes their ability to tolerate the procedure as well as the safety implications relating to metallic objects. The screening process will include:

- Verbally asking the subject the safety checklist questions, ensuring they understand the question and obtaining full answers, prompting for further details if necessary.

- A full explanation of Magnetic Resonance Imaging and the potential hazards within the Controlled Areas - what the scan involves, including duration, noise and their position within the magnet

The subject and Authorised staff member will sign the checklist. Screening must take place every time a subject attends an MR facility.

  - Volunteers

Healthy volunteer imaging as part of in house testing of sequences i.e.  Development work will be managed through the CRIF Approval Group.

Healthy volunteer imaging as part of a study cohort will be managed through UoG approval process.

Patient volunteer imaging will be managed through the standard NHS R&D Ethics approval process.

All volunteers should be consenting and fully informed using approved documentation.

Volunteers taking part in trials over a period of time should undertake safety screening at each visit.

Volunteers must not be under 16 years of age or known to be pregnant.

  - Incidental findings

At present the 7T facility is not used for clinical imaging therefore images will not be reviewed by a clinician.

If an obvious abnormality is detected by imaging staff then the volunteer would be referred for clinical imaging at 3T with appropriate onward referral thereafter. Refer to SOP 58.006 CRIF Reporting incidental findings.

  - Subject Clothing

All Subjects will be changed into appropriate clothing.

  - Subject Weight and Height

Currently Volunteer subjects may only be scanned with the 1Tx mode therefore all subjects should be weighed in the MR facility and their height measured before scanning so that when in 1Tx, the scanner is able to perform suitable accurate SAR calculations prior to execution of any imaging sequence and to assess cumulative energy dose.

  - Claustrophobia

Some subjects experience claustrophobia on entering the narrow bore of the MRI scanner.  To minimise anxieties the following can be implemented:

- Coaching by MR Operators, including careful introduction to the scanner.

- Use of mirrors to allow participants to see outside of the scanner and a call button which subjects can use at any time to communicate with the MR Operator or stop the scan.

- Suspending scans if subjects report discomfort or claustrophobia

  - Cable positioning

Poor positioning of subject and associated cables, leads and sensors, have been the cause of many burns reported to the MHRA.

Currently the 7T system is limited to head coils only with limited scope to cause burns. The coils and cables should be checked before every use for damage and do not use if damage seen. Report damage to manufacturer as soon as possible and complete fault log.

  - Acoustic Noise

All subjects and volunteers being scanned and personnel within the room during scanning should wear hearing protection.

  - Communication

All subjects should be spoken to as soon as MR Operator enters Control Room after positioning to check they are comfortable and can hear the MR Operator.

The MR Operator should then speak to the subject between every sequence to ask the subject how they are and an audible response should be obtained. If in any doubt the MR Operator must enter the Scanning Room to check the subject. The subject must also be told the duration of the next scan and any instructions that may be required.

  - Subject-Alert System/ Monitoring

All subjects are given the buzzer ball with an explanation of its intended use. If subject unable to use the buzzer ball then someone should remain in Scanning Room with subject.

The MR Operator must always maintain visual and audio contact with the subject. They must never leave the Control Room unless to enter the Magnet Room. This includes allowing access for Unauthorised Personnel into the Controlled Area.

  - Records

The participant screening checklist and consent form should be kept for a minimum of 5 years.

An imaging log will be maintained on the UoG server and CRIF database.

- Equipment

  - Faults and Maintenance

### 7.2.1  Responsibilities

Responsibility for performance of the Siemens Terra in the according to the parameters laid down in “Guidelines for Magnetic Resonance Diagnostic Equipment in Clinical Use” rests with the manufacturer/supplier, while the equipment is covered by the maintenance contract.

The equipment must not be interfered with or modified in any way, unless this is carried out (or specifically authorised) by the manufacturer.

The equipment will be covered by the Manufacturer’s contract

Under no circumstances must any unauthorised attempt be made by any Facility staff to repair the equipment or indeed perform any operation that may invalidate the Maintenance Contract.

7.2.2  Action to take

When a fault develops, the scanner should be restarted / shutdown and rebooted in the first instance.

Regardless of whether this solves the problem the manufacturer should be contacted by the MR Operator (and an MR Save Log performed on Siemens Scanner).  Give as much detail regarding the fault as possible including time, nature, examination / sequence being informed and any error messages. (Inform Siemens if you have performed MR Save Log.)

Usually you will be asked to leave your name and contact telephone number before an engineer will call you back. They will try and remotely diagnose the fault. Follow any instructions they give you, if in any doubt refer to the Supervising MR Operator.

### 7.2.3  Fault Log

A Fault Log on the UoG server will be maintained by the MR Operators, documenting any incident which occurs in relation to the MRI Equipment.

Accidents owing to serious defects in Magnetic Resonance Diagnostic Equipment must be reported to the Department of Health in accordance with NHS (Scotland) Circular 1991 (GEN) 24, “Reporting of Hazards and Potential Hazards”.

Any accident to subject or staff within the Facility must be reported on NHS GGC DATIX system in accordance with established NHSGGC policy.

### 7.2.4  Service Engineers

The manufacturer’s service engineers must report to the Supervising MR Operator in the Facility and prior to commencing maintenance work on the equipment. They must similarly report on completion of the maintenance and before leaving the Facility.

### 7.2.5  Quality Assurance Checks

The MR Operators/Physicists will be responsible for ensuring that quality assurance checks on equipment are carried out in accordance facility requirements. Long term QA data will be kept and routinely analysed by the Physicist.

### 7.2.6  Cryogen Refills

Cryogen refills will be carried out by the manufacturer within the context of the comprehensive maintenance service agreement.  No subject scanning may take place during any maintenance procedure on the cryogens

### 7.2.7  Ancillary Equipment

All ancillary equipment must be kept and stored in a safe and tidy manner.  In particular, damage to surface coils and leads should be avoided, as a damaged lead can inflict a subject burn.

## Emergency Procedures

In an emergency situation, access to the MR Environment must only be given under the direct supervision of an Authorised Person. All Authorised Persons must demonstrate a thorough knowledge of emergency procedures.  They must:

- Acquaint themselves with the sounds of the various alarm systems

- Familiarise themselves with the positions of the various emergency stop buttons and make sure that they understand the differences between them;

Magnetic Quench Buttons are the only way to turn off the magnetic field.  The financial consequences will result in several thousand pounds as the cryogens are lost.  Electrical systems are not shut down.  The couch brake is not released.  Please note that it will take up to a minute for the magnetic field to decrease to zero.

### Cardiac Arrest

Under no circumstances must a cardiac arrest crash team be allowed into the MR Environment.  In the event of a cardiac arrest, it is vital that the subject is moved to a point out with the MR Environment where resuscitation can safely begin.

Upon becoming aware of a cardiac arrest within the MR Environment, Facility staff will take the following action:

- Dial 2222: State: CARDIAC ARREST IN MRI, GROUND FLOOR, ICE building

- Move the subject as soon as possible to the adjacent anaesthetic area using the non-ferromagnetic trolley.

- Ensure the doors to the MR Environment are locked and the key placed in a draw in the Control Room.

- Oxygen should be administered as soon as possible using the nearest Crash Trolley.

- CPR should then be commenced if necessary.

- Resuscitation equipment must never be taken into the MR Environment.

### Fire

The main hazards to personnel from fire in the MR Environment are the effect of the magnetic force on ferromagnetic materials worn or carried by fire fighters and smoke produced by burning materials in the MR Environment such as plastic cladding foam etc. Therefore it is vital that as much material within the MR Environment is fire retardant, especially subject padding, bed covers etc.

#### Fire Extinguishing Equipment

All fire-fighting equipment located within the MR Controlled Access Area must be constructed of non-ferromagnetic materials and clearly marked as MR SAFE, in accordance with MHRA guidelines.

#### Action

It is imperative that all Facility staff and Authorised Persons are familiar with the appropriate action to take in the event of fire or a fire alarm going off in accordance with NHS GG&C procedure.

If smoke is smelt or electrical arcing is heard or if sparks are seen in the magnet room or in the computer cabinet then inform the MR Operator who will power down the scanner and get any subject out.

In the event of a fire developing within the MRI Facility the following action should be taken:

- Activate the fire alarm in the Waiting room - this will deactivate the secure doors

- Dial 2222, state: FIRE IN MRI, GROUND FLOOR, ICE building

- Evacuate the Scanner, lock door and secure key.

- Evacuate the Facility.

- If reasonably safe to do so, tackle the fire with the designated non-ferromagnetic extinguishers.

The Supervising Operator will remain to prevent the emergency services entering the MR Environment.

They may if absolutely necessary enter the Controlled Area with ferromagnetic equipment (but not if they have pace-makers) but they must not enter the MR Environment.

If the alarm is sounded for an incident out with MRI, the supervising MR Operator will act as Fire Warden in the evacuation of the Facility. They must ensure magnet door is locked and the key secure, that all staff and subjects are clear of the facility and go to the designated assembly point at the Neurology Building.

Failure to report the area clear will result in members of the fire service being sent to check the area and could delay the fighting of a real fire.

#### Quenching the Magnet

If the emergency service personnel must enter the MR Environment with ferromagnetic equipment or you feel that you cannot control them sufficiently to stop them doing this, or the fire is in the magnet room itself then you should hit the quench button.

N.B. Remember that after a quench it takes approximately 40 seconds for the field to fall to zero.  A quench can deplete oxygen within the Magnet Room to dangerously low levels.

The whole MR Facility must be evacuated. Ensure the emergency air extract system is activated.

Emergency service personnel must be informed of the oxygen deficiency and not allowed in unless they have their own portable air supplies.

Any Authorised Member of Staff has the authority to quench the magnet.

At all times, if there is a magnet quench, an attempt should be made to contact the MR Physicist/Responsible Person (by page and phone).

Evaluation of the MR environment after a quench is as follows:

- Check Oxygen level on Oxygen monitor, found on the far left hand corner of the control room – if the level is at least 20% then Magnet room is safe to enter.

- Check the Siemens display panel to determine status of Helium level and MAGSTOP – if both lights are illuminated the magnet field will have been reduced to zero.

However, do not enter the magnet room until final clearance is given by the MR Safety Expert / Responsible Person.

#### Authority

Once Fire Brigade personnel have arrived at the scene of the fire, responsibility for any decisions in respect of the well-being of personnel, buildings or equipment passes to the Fire Brigade Officer in charge, although in consultation with the MR Physicist or available Authorised Person.  The instructions of the Fire Officer in charge must be complied with.

It is imperative to check whether the Fire Officer in charge has been briefed before with the hazards of the Facility.  This may not be the case if the event occurs at a holiday time or otherwise busy period (additional information for the Fire Brigade is displayed at the entrance to the facility).

### Magnet Self-Quench

It is possible, though extremely unlikely, that the magnet may undergo a self-quench, either partial or total.  In general, quenching will be accompanied by,

- A loud bang

- Evolution of large quantities of cold gas

- Oxygen depletion in the atmosphere

- The risk of hypothermia and cold burns.

In the event of a partial quench, there may still be a significant residual static magnetic field with the associated risk of missile effect on ferromagnetic objects.

In the event of a self-quench:

- Turn off the scanner at the power source.

- Remove the subject IMMEDIATELY to a safe area OUTSIDE THE FACILITY.

- Ensure the emergency air extract system is activated.

- Evacuate the Facility of all personnel.

- Inform the Responsible Person, the MR Physicist and the Manufacturer.

The MR Environment should not be entered until the system has been examined by the manufacturer.  No further scanning should be carried out until the system has been declared fully operational by the manufacturer.

N.B. Remember that after a quench it takes approximately 40 seconds for the field to fall to zero.

### Flood

If water starts to leak into the magnet room, control room or computer cabinet, power down the system and turn isolation handles in the Equipment Room to completely isolate any remaining power.

Evacuate subjects and staff.

### Metal objects within the Magnet Bore

A ferromagnetic object entering the magnet bore is likely to cause damage both to the magnet structure and to personnel in its path.  If in spite of all precautions, such a situation should ever occur, the following procedures should be followed:

- No subject injury- remove the subject immediately to safety.  Do not attempt to remove the metallic object.

- Subject injured but not trapped - remove the subject immediately to safety out with the MR Environment and obtain medical assistance.  Do not attempt to remove the object.  Summon the Lead Radiographer and MR Physicist.  Attempting to remove a metal object may be hazardous as the changing force on the object will be difficult to predict.  This should not be attempted whilst anyone is in the magnet.

- Subject injured and trapped - medical assistance should be obtained immediately. If there is immediate danger to the subject’s life the magnet should be quenched immediately.  The subject should then be removed to safety out with the MR Environment for appropriate resuscitation measures and medical treatment. Following a quench, no one should enter the magnet room until clearance has been given by the MR Physicist / Responsible Person.

## Approved implants, medical devices and equipment for the 7T environment

Implants: At 7T, coils are local transmit/receive and no integrated body coil is used for transmission.

Given we operate in two modes, operators must follow the latest implant policy for each of these modes.

A full list implant’s authorised as safe to scan by the ICE Safety Committee can be found at- https://gla-my.sharepoint.com/:f:/g/personal/tracey_steedman_glasgow_ac_uk/EhJ5o3A1otJNpStDOBNiiv8B3ha7Ib578aSwkkkNCWnfiw?e=xKdUcG

Appendices

Appendix one: Plan of facility with magnet and 0.5 mT (5 Gauss line)

Appendix two: Safety labelling - definitions from ASTM international standard F2503-13

Appendix three: Safety Screening Checklists